Title: Sterilizer Door Pressure Interlock Fault

Document ID: DOC-2026-IPMS-STR-004

Category: Sterilizer Station Maintenance

Keywords: sterilizer door, interlock fault, steam pressure, door stuck, emergency exhaust, safety lock

Overview:

Occurs when the sterilizer door safety interlock sensor fails to disengage after steam discharge, preventing operators from opening the vessel door.

Symptoms:

- Door handle mechanical lock will not turn.

- Control panel displays "Interlock Active - Residual Pressure Detected".

- Pressure gauge needle oscillates near 0.1 to 0.3 bar.

Step-by-Step Resolution:

1. Ensure the main steam inlet valve is in the locked-closed (LOTO) position.

2. Manually open the auxiliary manual blow-down valve at the base of the sterilizer to vent trapped condensate and residual steam.

3. Wait until the auxiliary pressure dial reads exactly 0.0 bar (0 psi).

4. Inspect the mechanical interlock solenoid pin on the upper door hinge to ensure it has retracted completely.

5. If the solenoid pin remains extended, depress the manual safety release lever using a 10 mm brass rod.

6. Once the interlock disengages, rotate the handwheel counter-clockwise to swing the sterilizer door open.

7. Clean any oil palm fruit debris or sludge buildup around the door perimeter gasket before restarting the batch.